PROJECT FALCON-01

# Sensor Selection and
System Architecture

A defensible Phase 1 component baseline for prototype development, thesis documentation, calibration planning, and future marine deployment.

Revision 1.0

Prepared 31 August 2026
Adviser-aligned engineering baseline
Not yet a fabrication or deployment release

## 1. Executive decision

Project FALCON will estimate wave height from a submerged pressure time series. An ESP32 acquires and validates the measurements, then sends them through an approved cellular link to a shore Bay Station. The Bay Station stores the data, performs pressure-derived wave processing, runs the required prediction model, and serves the dashboard and alerts.

Core decision: the BNO085 is not required. Buoy motion remains visible in the dashboard, but the Phase 1 wave estimate is pressure-based. A separate low-complexity accelerometer may be added later only for tamper detection—not for the wave-height claim.

The exact prototype components are now identified where evidence supports them. Components that are not marine-qualified or have unresolved installation risks are explicitly marked as bench-only or proposed candidates.

## 2. Functional architecture

Marine environment ↓ Pressure · Wind · GPS · Water temperature ↓ Protected interfaces → ESP32 acquisition, validation and buffering ↓ Approved LTE/cellular modem → Mobile network ↓ Shore Bay Station ↓ Database → Pressure-derived wave estimate → AI prediction → Dashboard/alerts Battery + Solar → INA260 monitors → ESP32 health telemetry Enclosure contact (+ optional tamper accelerometer) → Security-state logic

The Orange Pi is not installed on the buoy under the current architecture. The shore Bay Station performs the heavier storage and processing tasks.

## 3. Status meanings

| Status | Meaning |
| --- | --- |
| Confirmed prototype | Exact part selected for bench and supervised prototype testing. |
| Recommended deployment candidate | Technically preferable for the marine role, but procurement, installation, calibration and field validation remain required. |
| Bench only | Useful for development but must not be presented as a final marine-deployment component. |
| Optional | Not required for the Phase 1 monitoring claim. |
| TBD | No exact product is approved; the item blocks final fabrication release. |

## 4. Core environmental sensors

| Function | Selected component | Interface / key data | Placement and validation |
| --- | --- | --- | --- |
| Pressure and wave-height input | Blue Robotics Bar02 R2, BR-100891 Confirmed prototype | 2.5–5.5 V supply; 2.5–3.6 V I2C logic. Verify address on the purchased unit. | Protected port at a documented submerged depth with no trapped air. Compare at known depths and record baseline, temperature and installation offset. |
| Position and geofence | Adafruit Ultimate GPS, PID 746 Confirmed prototype | 3.0–5.5 V board input; UART NMEA at 9600 baud default; 1–10 Hz. Use ~3 m nominal accuracy until site tests. | Clear sky view in upper assembly. Test cold/warm starts and stationary scatter. Geofence radius must exceed observed scatter. |
| Wind speed and direction | SparkFun Weather Meter, SEN-15901 Confirmed prototype | Reed pulse for speed; resistor-network vane through protected 3.3 V ADC. 1 pulse/s corresponds to 1.492 mph. | Unobstructed mast position, aligned to north. Compare with reference instruments and inspect corrosion, contacts and bearings after salt exposure. |
| Water temperature | Blue Robotics Celsius R2, BR-100317 Recommended deployment candidate | I2C; manufacturer states ±0.1 °C. Verify voltage, address and harness before PCB freeze. | Correct sealed bulkhead installation; rear electronics remain dry. Perform three-point bath comparison and leak test. |
| Water-temperature alternative | Adafruit waterproof DS18B20, PID 381 Bench only | 3.0–5.5 V, 1-Wire, 4.7 kΩ pull-up; ±0.5 °C from −10 to 85 °C. | Short supervised testing only. The manufacturer excludes salt water, corrosive environments, long-term use and high pressure. |

Pressure-sensor deployment limitation: the Bar02 sensing face is waterproof, but its rear is not. The manufacturer also requires the gel element to dry for at least two hours each day. Continuous long-duration deployment therefore needs a documented service plan or a differently qualified pressure sensor.

## 5. Health and security sensing

| Function | Selected component | Interface / key data | Placement and validation |
| --- | --- | --- | --- |
| Battery monitor | Adafruit INA260, PID 4226 Confirmed prototype | I2C `0x40`; up to 36 V bus and 15 A continuous on breakout; integrated 2 mΩ shunt. | Protected battery branch. Compare voltage/current with calibrated DMM and load at three operating points; inspect connector heating. |
| Solar monitor | Adafruit INA260, PID 4226 Confirmed prototype | I2C `0x41`; same limits as battery monitor. | Document measurement point and current direction. Validate during low, normal and high charge conditions. |
| Enclosure temperature | Adafruit MCP9808, PID 1782 Recommended selection | I2C `0x18`; 2.7–5.5 V; typical ±0.25 °C. | Representative enclosure air, away from modem/converters/heatsinks. Compare after thermal stabilization. |
| Enclosure-open state | Adafruit magnetic contact, PID 375 Recommended prototype selection | Normally open contact; filtered/debounced GPIO; fail-safe wiring preferred. | Inside dry enclosure or sealed interface. Validate alignment, cable faults, vibration and false alarms. |
| Tamper aid | Adafruit LIS3DH, PID 2809 Optional security candidate | I2C at `0x19` or SPI; ±2/4/8/16 g. | Rigid tray mount. Not a wave sensor. Thresholds require supervised normal-wave and tamper tests. |
| Audible alarm | Exact buzzer TBD | GPIO through transistor/MOSFET driver; flyback protection if inductive. | Voltage, current, sound level, duty cycle and environmental rating must be selected and tested. |

## 6. Communication candidate

Waveshare SIM7600G-H 4G HAT is a proposed communication candidate, not a sensor and not yet frozen. It uses a separate regulated 5 V branch and supports USB or level-compatible UART. Approval requires a Philippine carrier-band and site-coverage survey, antenna review, measured registration/transmit peaks, brownout recovery, reconnect behavior, offline buffering, operating-temperature checks and a complete power budget.

## 7. Interface allocation

| Interface | Device | Allocation | Engineering note |
| --- | --- | --- | --- |
| I2C 3.3 V | Bar02 R2 | Expected `0x76`; verify | Run an I2C scan on purchased hardware. |
| I2C 3.3 V | INA260 battery / solar | `0x40` / `0x41` | Record address straps and current direction. |
| I2C 3.3 V | MCP9808 | `0x18` | Keep placement thermally representative. |
| I2C 3.3 V | ADS1115 wind direction | `0x48` | Use a 3.3 V divider; verify ADC range. |
| I2C/SPI optional | LIS3DH | `0x19` or SPI | Avoid collision with MCP9808 at `0x18`. |
| UART | GPS | Dedicated UART | Cross TX/RX; verify logic voltage and approved GPIOs. |
| UART or USB | LTE modem | Dedicated link | Do not block GPS acquisition; freeze after system tests. |
| GPIO pulse | Anemometer | Interrupt-capable input | Protection, pull-up, debounce and pulse-rate tests required. |
| GPIO | Enclosure contact | Filtered input | Prefer cable-fault-detecting behavior. |

Exact GPIO assignments remain governed by the verified schematic and PINOUT.md. This architecture table does not override physical pin validation.

## 8. Removed or excluded items

- BNO085 sensor fusion is removed from the required Phase 1 design.

- Load-cell/HX711 anchor-chain tension sensing is removed.

- Salinity/conductivity is outside Phase 1 scope.

- The Orange Pi is not installed on the buoy.

- The optional LIS3DH is for tamper evidence only, not wave-height estimation.

## 9. Calibration and acceptance

### Pressure

- Inspect gel, connector, seal and cable.

- Record raw pressure in air and at multiple known water depths.

- Compare with a traceable reference and record temperature.

- Check for trapped air, leaks, drift and spikes.

- Freeze the installation offset and algorithm version.

- Label the output “Estimated wave height” until method validation is complete.

### Temperature

Use a stirred bath or stable chamber at no fewer than three temperatures. Allow stabilization, compare with a traceable thermometer, calculate offset/repeatability, and separately leak-test the deployed probe and bulkhead.

### Wind

Verify mechanical north alignment, every direction code, several steady speeds, low-speed startup, contact bounce, cable movement and salt contamination.

### GPS and security

Measure stationary GPS scatter at the site before setting a geofence. Exercise normal waves, maintenance, enclosure opening and simulated tampering. Ordinary motion must not create a theft alarm.

### Power and thermal health

Compare both INA260 channels with a calibrated DMM and controlled load. Confirm polarity, address, connector heating and error. Thermal-test MCP9808 placement while the modem transmits and converters carry normal load.

## 10. Procurement and release checklist

Record the following before purchase or PCB fabrication:

- manufacturer, exact SKU, revision and supplier;

- official datasheet and archived copy;

- measured dimensions, connector and pinout;

- supply voltage, logic level, normal and peak current;

- environmental and ingress limitations;

- footprint and 3D-model match;

- calibration equipment and acceptance limits;

- spares, availability and lead time.

Fabrication release remains blocked until the water-temperature installation, buzzer, LTE interface, connector variants, footprints, cable glands, current budget and enclosure integration are physically verified. A schematic render or 3D model alone is not validation.

## 11. Thesis-safe system statement

The Phase 1 Project FALCON monitoring system uses a Bar02 R2 pressure sensor, GPS, wind speed/direction sensing, battery and solar electrical monitoring, enclosure temperature, and a marine-capable water-temperature candidate. The ESP32 acquires and validates sensor data before cellular transmission to the shore Bay Station. The station processes pressure history to estimate wave height and support AI prediction. GPS geofencing and an enclosure contact provide security information. Deployment readiness will be claimed only after calibration, ingress, power, LTE, environmental and integration tests are completed.

## 12. Primary manufacturer references

- Blue Robotics. Bar high-resolution pressure sensors.

- Blue Robotics. Bar sensor technical guide.

- Blue Robotics. Celsius fast-response temperature sensor.

- Adafruit. Ultimate GPS, PID 746; GlobalTop PA1616S datasheet.

- SparkFun. Weather Meter Kit, SEN-15901.

- Adafruit. Waterproof DS18B20, PID 381.

- Adafruit. INA260, PID 4226.

- Adafruit. MCP9808, PID 1782.

- Adafruit. LIS3DH, PID 2809; STMicroelectronics datasheet.

- Adafruit. Magnetic contact switch, PID 375.

- Waveshare. SIM7600G-H 4G HAT.